Supplementary Material for the VALIDATE-HF Analytical Framework

Appendix A: Important Considerations in DHT Data Collection

Standard practices for actigraphy measurements are as important as the reliability assessment itself. Here we present additional information on how to prepare and design the most effective validation method for DHTs.

A.1	Practical Considerations

To collect accurate and reliable data for actigraphy measurements, researchers should consider the limitations of their cohorts carefully, such as physical limitations, DHT placement, duration of the measurement, among others. Validation of actigraphy DHTs may require manual measurement (such as 6MWD or step count) or other validated DHTs (sleep quality, or activity rate). The first step in the validation is ensuring that the consistency in the measurement setting and data quality.

Sampling actigraphy data, in addition to referring to specific guidelines for collecting data from the unvalidated DHT, also encapsulates data collection practices for the gold standard. Most common issue in actigraphy measurements are inter-rater and intra-rater variability, as well as changes due to the environment (time of the day, location, etc.).

Furthermore, data acquisition practices should report the sampling rate of the DHT as well as the gold standard to ensure no aliasing effect or under sampling occurs during feature extraction. For example, activity counts for a sampling rate of 30 Hz and 100 Hz for the same device can have a bias as large as 1,000 –3,000 more counts per minute, accounting for more than 10% difference 18. In addition, outliers and preprocessing the data will minimize the effect of rater-based errors, giving a clearer picture of the DHT reliability.

A.2	The Difference Between Agreement and Correlation

As mentioned in the framework, correlation measures the strength of the linear relationship between two variables, whereas agreement measures the consistency between them, meaning how well they match.

Two variables could be strongly linearly related, hence highly correlated, but not agreeing. For example, temperatures in °C and °F are highly correlated as they are linearly proportional but do not agree, since 0°C is not the same temperature as 0°F, as illustrated by Appendix Figure 1. This applies equally to °C and K; although they employ the same scale, the 273-degree shift results in a significant discrepancy, and thus, no agreement. Hence, considering the use case study aims to evaluate the agreement between manual and digital 6MWD measurements, a high correlation index (ρs) is insufficient to demonstrate the Apple Watch's accuracy. However, correlation is still an important measure to consider, as it provides a sense of the dataset before assessing for agreement.

The relationships between °C and °F and °C and K demonstrates the nuance between correlation and agreement. All variables provide measurements of the same metric, temperature, but vary according to different scales, as illustrated by the shifts of the °F (red) and K (blue) lines from the 45° (y=x) perfect agreement line (black). Hence, they are perfectly correlated (ρs=1), but do not agree, as demonstrated by the extreme CCC and MAPE values, which reflect how 1°C is very different from 1°F or 1 K in temperature.

To extend this logic to DHT validation, comparing the measurements of a digital device with the gold standard (e.g., manual measurements) includes assessing the agreement between measurements yielded by both methods. For example, when assessing the reliability of a smartwatch in measuring walking distance, we want to know whether a smartwatch’s measured distance is equal to (i.e., agrees with) the corresponding manual measurements; hence, assessing for correlation only constitutes an incomplete methodological approach.

A.3	Individual-Level and Systematic-Level Reliability

Evaluating the reliability of a DHT implies assessing its fitness-for-purpose; however, given that DHTs can serve diverse purposes, the statistical methodology employed in their analytical validation must be tailored according to the context of use of the device of interest.

Thus, we divided the metrics into two different categories, depending on whether they assess reliability on an individual or a systematic level. On one hand, individual-level reliability, assessed by CCC, ICC, and BA LoAs, pertains to the reproducibility of measurements on an individual level, meaning how well each individual’s corresponding measurements (device vs gold standard) agree. On the other hand, systematic-level reliability, assessed through MAPE and Bland-Altman analysis focused on bias, relates to agreement within measurements at the level of the tested group as a whole.

Consider the following example: two different thermometers are used to measure room temperature every hour, over the course of a week. When comparing the measurements of two different thermometers to measure room temperature, a high individual-level reliability would indicate that both thermometers yield similar measurements on each individual hourly reading, whereas a high systematic-level reliability would be indicated by both thermometers yielding similar mean temperatures each day without necessarily providing identical readings every hour.

A.4 Regulatory Considerations

The FDA guidance on integration of DHTs into clinical investigations serves as a crucial compass for industry professionals, investigators, and stakeholders alike, offering invaluable insights into the adept utilization of DHTs for remote data acquisition within clinical investigations.1

Currently, the FDA has approved 62 artificial intelligence-driven DHTs in the field of cardiology that either measure functional status or improve HF treatment trackingpatients.2 However, to date, there has been little systematic evaluation of actigraphy-based endpoints collected by DHT against recognized functional endpoints.

The FDA's guidance emphasizes that the application of DHTs must be "fit-for-purpose," signifying that the level of validation accompanying these tools should be commensurate with their intended use and interpretability in clinical investigation. Key considerations encompass data accuracy, reliability, and security. DHTs enlisted for clinical trials must be meticulously selected, considering a range of factors, including the nature of data to be gathered, the target patient population, and any relevant regulatory prerequisites. The rationale for selecting a specific DHT and its intended use should be substantiated, thereby ensuring alignment with the study's objectives.

Clinical endpoints must be rigorously justified and precisely defined, encompassing the timing and tools employed for assessments. DHTs may replace established functional endpoints, necessitating validation to guarantee data reliability. The statistical methodologies used for analyzing data collected via DHTs should be clearly and pre-emptively defined, with consideration given to potential intercurrent events that may impact data integrity.

The rise of decentralized clinical trials (DCTs) heralds a transformative shift in clinical trial methodology. One such trial on SGLT2 inhibitors in HF highlights the importance of DCTs in novel therapies.3,4 These trials empower participants to receive treatments, undergo assessments, and provide data from the comfort of their homes or local healthcare facilities. The use of DHTs in DCTs offers several advantages, including increased patient access, enhanced convenience, personalized engagement, efficient data collection, and cost reduction.

Despite the evident benefits of DCTs, adherence to applicable regulations and guidelines remains paramount. Study design, patient recruitment, data collection, monitoring, and ethical considerations must align with regulatory standards to safeguard human subjects, ensure patient safety, and maintain data quality while upholding regulatory compliance in clinical trials. A summary of the interpretation and considerations are given in Appendix Figure 2.

Appendix B: Computation and Interpretation of the Relevant Metrics

B.1	Computing the Relevant Metrics

The description of the relevant reliability coefficients and metrics mentioned in the framework, as well as the formulas used to compute these metrics, are outlined in Appendix Table 1 and Appendix Table 2, respectively. As illustrated in Appendix Table 1, we divided the metrics into two different categories: individual reliability and systematic reliability-testing metrics. Individual-level reliability pertains to the reproducibility of measurements on an individual level, whereas systematic-level reliability relates to agreement in measurements at the level of the tested group as a whole. The distinction between individual and systematic-level reliability is described in Appendix A.4. As there are 10 different ways to compute the ICC, its formula selection process is described below.

B.2	ICC Formula Selection

There are 10 different ways to compute the ICC, and selection of the appropriate formula depends on the study design. ICC formula selection is a three-step process, consisting of model, type, and definition selection.

First, model selection pertains to the measurement methods and consists of choosing between three options. A one-way random effects model is used in a study involving multiple different devices randomly assigned to participants (e.g., some participants wearing the Apple Watch, other wearing the Fitbit, all compared to the manual gold standard). A two-way random-effects model is indicated when the aim is to generalize the results to all devices similar to the investigated one (e.g., using an Apple Watch in a study assessing the reliability of commercial smartwatches in general in measuring walking distance). In contrast, a two-way mixed-effects model should be selected when the devices used in the study are the only devices of interest (e.g., using an Apple Watch in a study specifically assessing its reliability in measuring walking distance).

After selecting the model, the next step is choosing between two types of ICC: single rater and multiple raters. The single rater method consists of comparing the investigated device’s measurements to a single criterion method (e.g., comparing smartwatch measurements to a manual gold standard), while the multiple raters method consists of comparing the device’s measurements to the mean of multiple raters, or methods of measurements (e.g., investigating the Apple Watch, Fitbit, and Withings Go watches and comparing the Apple Watch to the mean of measurements of all watches).

The final step consists of selecting the relevant ICC definition. The absolute agreement ICC tests for perfect agreement between corresponding measurements of different methods (i.e., y = x), whereas the consistency ICC is used to assess the consistency of the systematic bias between methods (i.e., y = x + c, where c is the bias).

The consistency ICC is a useful metric in cases of unsatisfactory agreement but constant systematic bias (e.g., a constant overestimation of ≈25m in smartwatch walking distance measurements) to quantify the consistency of the error between measurements.

Of note, although the ICC is recommended for quantification of correlation and agreement between three or more devices or measurement methods, it can still be used to compare two methods of measurements. ICC and CCC usually have similar values when only 2 raters are present in the dataset.5 Besides, the ICC can be used to measure repeatability using the Spearman-Brown formula, paired with Bland-Altman analysis conducted with the mean of measurements throughout time (e.g. repeating digital 6MWD measurements daily, computing the mean of all measurements and comparing it to the original mean of one-day measurements)

Appendix Table 2 provides the appropriate formula for each combination of the three components of ICC selection.

B.3	Interpretation of Metrics

We propose a guideline to interpret the results of the computed metrics in Appendix Table 4.

Appendix C: Use Case Dataset and Discussion

C.1	Use Case Demographics

The dataset of the pilot study used as an example in the manuscript is from a prospective case-control study involving 1584 male and female participants aged between 35 and 79 years who underwent the 6-Minute-Walk-Test (6MWT) indoors with an Apple Watch Series 3 (Apple Inc., CA, USA) using its Indoor Walk mode. All participants were enrolled via the Courtois Cardiovascular Signature (CCVS) Program, a clinical research program aiming to personalize the management of cardiovascular health using analytic approaches centered around digital health and artificial intelligence, genetics, and microbiomics, as well as environmental factors to optimize cardiovascular patient outcomes. The program involves 4000 participants in a longitudinal prospective case-control study over a 10-year period, through which information on cardiovascular events, diagnoses, and medication is collected in the CCVS biorepository. The participants for this study were approached first by their physician at their cardiovascular clinic visit and their informed consent to the program was obtained.

Participants were separated into three cohorts (Appendix Table 5): patients with heart failure (HF, n=86), patients with other cardiovascular conditions (CVD non-HF, n=172), and healthy participants (healthy, n=1326). Outliers were excluded based on the Q1/Q3 ± 1.5 IQR criterion, rendering the final sample sizes to be: n=74 (HF), n=161 (CVD non-HF), and n=1263 (healthy). Heart failure was defined as clinical signs and/or symptoms caused by a structural and/or functional cardiac abnormality, as defined by Bozkurt et al.,6 diagnosed at the cardiovascular clinic, as well as natriuretic Peptide (NT-proBNP) levels above 125 pg/L or objective evidence of cardiogenic pulmonary or systemic congestion. The HF cohort comprised patients with well compensated heart failure, who received regular cardiology care and had their quality of life optimized in spite of their condition. Additional exclusion criteria consisted of pregnancy, general MRI contraindications, and no understanding of English and French.

Based on American Thoracic Society Guidelines, participants were asked to complete a 6MWT as part of their standard of care. The test includes asking participants to walk back and forth between two cones along a corridor of 28 m lap length for six minutes wearing the Apple Watch on their dominant hand while the time and walking distance were being manually recorded. Patients were allowed to slow down, stop or rest during the task if necessary. Fatigue and shortness of breath (SOB) severity were collected using the modified Borg scale, along with oxygen saturation (SpO2), heart rate (HR), and systolic and diastolic blood pressure (SBP, DBP). Participant height was entered in the phone paired to the Apple Watch prior to the 6MWT for calibration.

C.2	Correlation and Agreement Plots

This section includes the plots (Appendix Figure 3 and Appendix Figure 4) illustrating correlation and agreement, reflecting ρs and BA respectively. The correlation between digital and manual measurements of 6MWD is illustrated in the Appendix Figure 3 for each cohort. Manual 6WMT-measured distance is represented on the x axis, while the Apple Watch’s digital 6MWD distance is represented on the y axis. Each scatter plot displays the linear regression line and the corresponding ρs value.

C.3	Bland-Altman Analysis

Bland-Altman analysis is used to evaluate the systematic bias between digital and manual measurements by comparing it to the relative MCID of 10% for HF patients in the case of the Apple Watch study. In other words, digital measurements exhibiting a discrepancy exceeding 10% relative to their manual counterpart were categorized as clinically unreliable Apple Watch measurements.

The Bland-Altman plots for the HF, CVD non-HF, and healthy cohorts provided in Appendix Figure 4 highlight systematic overestimation of 6MWD by the Apple Watch, as illustrated by the mean bias line being above the perfect agreement line (y = 0%) in each cohort. Bland-Altman analysis yielded mean biases (CIs) of +5.9% (3.7, 8.1), +6.4% (4.8, 8.0), and +6.3% (6.8, 5.8) in HF, CVD non-HF, and healthy cohorts respectively. Hence, in every group, the mean bias and its 95% CI fall within the clinical threshold of ± 10%.

Bland-Altman plots also provide the opportunity to assess the spread of the data beyond simply interpreting the LoAs; looking for patterns in the spread of the dataset allows to draw additional conclusions on the tested device. For example, increasing absolute differences between digital and manual measurements (y axis) with incrementing average values (x axis) on a standard BA plot, but relative differences (y axis, %) staying constant on the corresponding relative BA plot could indicate that the device accumulates consistent small errors in measurements throughout the walk, thereby making agreement and reliability a function of the traveled distance (6MWD). Sevrukov et al. provide an example where the absolute difference between measurements appears to be an increasing function of the average.7 Although the lack of such a pattern suggests a different explanation for the suboptimal accuracy of the Apple Watch, it is always important to appreciate for patterns in the spread of the dataset.

Of note, Bland-Altman analysis can also be used to evaluate the replicability of multiple (i.e., two or more) DHT measurements, provided the measurements are taken within a short enough timeframe to avoid results being confounded by changes in the patient’s disease state. As an example, one could record 6MWD measurements on four consecutive days and assess whether the measurements from different days are in agreement with each other to evaluate the replicability of the results.27

C.4	Interpreting Metrics

First, the Spearman correlation coefficient (ρs) was computed to quantify the strength of the relationship between Apple Watch digital measurements and manual 6MWT coordinator measurements of distance. The ρs values indicated moderately strong correlation between the manual and digital 6MWD across all cohorts. Nevertheless, the CCC indicated poor levels of agreement for every group. The discrepancy between CCC and ρs values in each cohort highlights the difference in their respective functions and the nuance between correlation and agreement, showing that strong correlation is not synonym with satisfactory agreement.

Next, the ICC, giving insight into measurement replicability and capturing both correlation and agreement into a single metric, yielded moderate reliability levels using the two-way mixed-effects, single-rater, absolute agreement ICC.

The MAPE was then computed to quantify forecasting accuracy and interpreted using Nelson et al.’s threshold,8 also used by Boudreaux et al.9 and Vetrovsky et al.10 in the DHT literature, setting a MAPE ≤ 10% as the validity criterion. Of note, other thresholds exist in the literature, such as Fokkema et al.’s 5% MAPE validity threshold,11 but are used less frequently. In the context of this study, despite MAPE values falling below the 10% threshold in every cohort, indicating satisfactory systematic agreement, the values were marginal and approaching the 10% cut-off point (HF: 9.2%, CVD: 9.8%, and Healthy: 9.5%), which may warrant critical appraisal and careful interpretation.

C.5	MCID Selection

Although DHT reliability assessment is based on a group of diverse metrics constituting a comprehensive statistical approach, Bland-Altman analysis, which constitutes the main component of agreement assessment, must ultimately yield a dichotomous output on the systematic reliability of a device in clinical research (i.e., reliable or not reliable). This requires setting a clinically relevant threshold appropriate to the specific test and patient population is necessary for the assessment of systematic reliability: the minimal clinically important difference (MCID).

This use case investigates the reliability of a device in measuring 6MWD in HF patients, so we selected a MCID appropriate to the task. Of note, the significance of the absolute MCID (in m) is dependent on patient 6MWD, as a 50 m 6MWD difference is not as significant in a patient with a 600 m 6MWD as it is in a patient with a 250 m 6MWD because differences are more significant to HF patients with advanced disease progression and severe symptoms. Hence, the relative MCID (in %), coupled with the relative difference (percentage) variation of the Bland-Altman plot, 12 represents a more universal accuracy threshold independent of disease severity, thus providing more relevant insight on device accuracy. Therefore, we set 10% as the clinically relevant threshold for the 6MWD, as it has been proposed as the relevant MCID by the Canadian Heart Failure Society ,6 as well as the HFC-ARC Panel in the Journal of the American College of Cardiology (JACC).13

References

1.	FDA. Digital Health Technologies for Remote Data Acquisition in Clinical Investigations Guidance for Industry, Investigators, and Other Stakeholders. Published online 2021.

2.	FDA. Artificial Intelligence and Machine Learning (AI/ML)-Enabled Medical Devices. FDA. Published online October 5, 2022. Accessed October 3, 2023. https://www.fda.gov/medical-devices/software-medical-device-samd/artificial-intelligence-and-machine-learning-aiml-enabled-medical-devices

3.	Spertus JA, Birmingham MC, Butler J, et al. Novel Trial Design: CHIEF-HF. Circ Heart Fail. 2021;14(3):e007767. doi:10.1161/CIRCHEARTFAILURE.120.007767

4.	Spertus JA, Birmingham MC, Nassif M, et al. The SGLT2 inhibitor canagliflozin in heart failure: the CHIEF-HF remote, patient-centered randomized trial. Nat Med. 2022;28(4):809-813. doi:10.1038/s41591-022-01703-8

5.	Chen CC, Barnhart HX. Comparison of ICC and CCC for assessing agreement for data without and with replications. Computational Statistics & Data Analysis. 2008;53(2):554-564. doi:10.1016/j.csda.2008.09.026

6.	Bozkurt B, Coats AJS, Tsutsui H, et al. Universal definition and classification of heart failure: a report of the Heart Failure Society of America, Heart Failure Association of the European Society of Cardiology, Japanese Heart Failure Society and Writing Committee of the Universal Definition of Heart Failure: Endorsed by the Canadian Heart Failure Society, Heart Failure Association of India, Cardiac Society of Australia and New Zealand, and Chinese Heart Failure Association. Eur J Heart Fail. 2021;23(3):352-380. doi:10.1002/ejhf.2115

7.	Sevrukov AB, Bland JM, Kondos GT. Serial electron beam CT measurements of coronary artery calcium: Has your patient’s calcium score actually changed? AJR Am J Roentgenol. 2005;185(6):1546-1553. doi:10.2214/AJR.04.1589

8.	Nelson MB, Kaminsky LA, Dickin DC, Montoye AHK. Validity of Consumer-Based Physical Activity Monitors for Specific Activity Types. Medicine & Science in Sports & Exercise. 2016;48(8):1619-1628. doi:10.1249/MSS.0000000000000933

9.	Boudreaux BD, Hebert EP, Hollander DB, et al. Validity of Wearable Activity Monitors during Cycling and Resistance Exercise. Medicine & Science in Sports & Exercise. 2018;50(3):624-633. doi:10.1249/MSS.0000000000001471

10.	Vetrovsky T, Siranec M, Marencakova J. Validity of six consumer-level activity monitors for measuring steps in patients with chronic heart failure. PloS One. 2019;14(9). doi:10.1371/journal.pone.0222569

11.	Fokkema T, Kooiman TJM, Krijnen WP, Schans VANDER, CP DEG, M. Reliability and Validity of Ten Consumer Activity Trackers Depend on Walking Speed. Med Sci Sports Exerc. 2017;49(4):793-800. doi:10.1249/MSS.0000000000001146

12.	Giavarina D. Understanding Bland Altman analysis. Biochem Med (Zagreb). 2015;25(2):141-151. doi:10.11613/BM.2015.015

13.	Psotka MA, Abraham WT, Fiuzat M, et al. Functional and Symptomatic Clinical Trial Endpoints: The HFC-ARC Scientific Expert Panel. JACC Heart Fail. 2022;10(12):889-901. doi:10.1016/j.jchf.2022.09.012